{"text_format":"2","workbook":"xlsx","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":30,"n_cols":5,"rows_kept":30,"cut":false,"hidden":false,"merged":[[0,0,0,4]]}
[" ITB-163/17 - SUPPLY AND INSTALLATION OF BLOOD BANK EQUIPMENT 4 LOTS IN:\nLot3:  Blood Bank  Al Khansaa Hospital - Ninewa Governorate"]
["Item NO.","Detailed Description","Quantity  ","Unit Price USD","Total Price\nUSD"]
[1,"\nHemoglobin meter \nPURPOSE: For hemoglobin estimation\nLight source: LED, adjusted to specified wavelength, bandwidth: ± 2nm\nDetector: Silicone photodiode other should be specified \nMeasuring: Transmission/Absorption\nPhotometric accuracy: ± 0.002 \nCuvette types: should accept different type of cuvettes \nData: 1000 or more determinations including date and time\nDisplay: hemoglobin level in g/dl or g/l, date and time\nMeasuring range:  0 - 30g/dl Hb\nDilution ratio: should be stated \nLight source : Bright Intensity LED, or better  \nSample volume: minimum sample volume should be stated \nCalibration: yes\nQuality control: yes, with control cuvette\nSample material :Venus ,arterial or capillary\nEnvironmental requirements: the equipment suitable for working in the climate conditions of Iraq in terms of temperature & humidity\nOpen system : Yes \n     Power supply: 230V,50Hz, \nSample compartment: should be prepared for the use with cuvettes, integrated cover for the protection of the sample compartment against ambient light and dirt\n",5]
[2,"\nBlood bag floor standing centrifuge 6-12 bag\nApplication: separate suspended particles of blood bags different densities separated blood components.\n Rotational speed:\n• Max rpm: (1000 – 7000 ).\n• Max rcf, g: (≥4000).\n• No.of  blood bag for each cycle: ≥ 12 bags \n• Speed adjustment: microprocessor controlled.\n Height:≥ 90cm\n Tachometer accuracy (rpm) : (≥1).\n Rotors available:  horizontal .\n Temperature range: (-20~+60).\n Timer: yes.\n Braking (electric type): yes.\n Brush ∕ brushless: brushless.\n Safety feature:\n• Lid interlock: yes.\n• Biohazard containment: yes\n Alert indicator: imbalance, end of run, overspeed, lid open & others.\n Display type: digital LCD .\n Programmable: yes.\n Noise level: (≤70 db).\n Input power: 230V,50Hz,  single phase.\n",2]
[3," Can tube centrifuge 5ml capacity\n\nSPECIFICATION\n• Spins 3ml-5ml test tubes at 5000 rpm \n• Minimum 16 tubes\n• Easy cleanout bowl\n• 0 - 5000 RPM \n• Digital programmable control and readout \n• Variable speed and time settings \n• Lid safety lock \n• high g-forces for fast separation \n• Rubber feet to prevent slipping \n• Impact-resistant lid \n• Zero-rpm interlocking lid \n• Wide mouth for easy clean-up \n• Metal rotor and tube shields\n• Brushless motor \n• Angled and (swinging )tube holder as alternative\n• Fast breaking\n• Programmable time and speed \n• Rubber feet to prevent slipping \n• POWER REQUIRMENTS : 230V,50Hz, ",4]
[4,"Bench type 12 plate centrifuge\nSPECIFICATION\n• Spins 3ml-15ml test tubes at 5000 rpm \n• Minimum 16 tubes\n• Easy cleanout bowl\n• 0 - 5000 RPM \n• Digital programmable control and readout \n• Variable speed and time settings \n• Lid safety lock \n• high g-forces for fast separation \n• Rubber feet to prevent slipping \n• Impact-resistant lid \n• Zero-rpm interlocking lid \n• Wide mouth for easy clean-up \n• Metal rotor and tube shields\n• Brushless motor \n• Angled and (swinging )tube holder as alternative\n• Fast breaking\n• Programmable time and speed \n• Rubber feet to prevent slipping \n• POWER REQUIRMENTS : 230V,50Hz, ",2]
[5,"Cell washed centrifuge \nMax. Speed                  4000rpm\nMax. RCF                       1800xg\nRBC Swing Rotor        12 x 7ml, 4000rpm, RCF 1800xg\nWBC Swing Rotor      12 x 2ml, 4000rpm, RCF 1800xg\n                                          12 x 1.5ml, 4000rpm, RCF 1800xg\n                                          12 x 0.5ml, 4000rpm, RCF 1800xg\nTimer                              Programmable\nMotor                             Brushless motor\nPower                             230V,50Hz, , 3A\nWeight                           22KG\nDimension                  430x346x260mm (L x W x H)",6]
[6,"Plasma extractor\nAutomatic Plasma Expresser\n• Purpose: separation of plasma from blood components of the blood bag\n• Inner & Outer  Body : Complete in Stainless Steel  \n• Pressure Plate : Acrylic\n• Clamping  : motor activated \n• Tube sealer with sensor \n• Type Sensor : infra-red \n• Operation : Automatic Controlled by PID (Programmable) microprocessor based controlled .\n• Highly sensitive sensor for instant sensing between transfers of blood components.\n• Alarm : audio visual & stop separation if other components transferred.\n• Should be compatible with different types/designs of bags\n• power requirement :  230V,50Hz, \n",4]
[7,"\nAphaeresis blood cell separator\n Application : (1) collection of blood products (i.e., donor collection) \n                              (2) exchange (removal and replacement) of substances \n                             (3) therapeutic removal of substances, or a combination of these     \n                              applications\n\n Type system  : closed system \n Separation method :built in centrifugation or filtration \n Centrifuge ,rpm : ≥ 8000  for centrifugation separation method\n Flow type : continuous \n venous access :  single \n Inlet rate, ml/min : range should cover the values 10- 150 .\n Extracorporeal volume, ml: minimal volume preferred.\n Anticoagulant ratio: varies \n apheresis units should allow the anticoagulant ratio and replacement fluid balance to be set manually\n programs  : yes \n monitors/alarms : yes , to be specified \n data management: yes \n battery backup to complete the cycle.Max. 1 hr.\n power requirements: 230V,50Hz, three pin g type plug , medical approved power supply board \n Environmental requirements: the equipment suitable for working in the climate conditions of Iraq in terms of temperature & humidity\n",1]
[8,"Tube sealer \n Blood bag tube sealer single\nPurpose: to seal blood bags tubes \n-Heavy duty for continuous sealing \n- Sealing Capacity:≥ 200/ day   \n- High-quality sealing \nAdjustable sealing time with range from 0.3ses to 6sec or less for permanent sealing\n-Power: 50 Ohm maximum or better\n-Frequency: ≥ 40.0MHz\n-Electric shock protections \n- single segment sealing \n-Input Power: 230V,50Hz, \n-Environmental requirements: the equipment suitable for working in the climate conditions of Iraq in terms of temperature & humidity\n",5]
[9,"Blood mixer \nRoller Size 4 Nos. x 230mm Lg\nRPM: 30 RPM/ 36 RPM\nMixing Time:15 - 20 Minutes\nCapacity Vacuum / Non vacuum Tubes 6 Nos(of 4 ml) or Bulb of 5ml x 15 Nos.\nWeight: 2.0 Kg\nPower Input: 230V,50Hz, \n",5]
[10,"\nElectric Blood Donation Chair\nPurpose: are especially made for the comfort of the blood-donor in the blood center and hospitals.\nType: motorized(electrical)\nFunctions movement :up/down, back section turn, and leg section turn.\ncastors :NO\nmaterial :high quality(strength).\nEasy to clean seamless cushions, Antimicrobial fabric\nAllows right or left-sided procedures\n• Chair dimension : ≥1900mm (curve length) x 550 mm\n• Arm board swing out : 0~≥ 45o\n• Arm board height adjustment : ≥ 70mm\n• Arm board slide ≥80 mm\n• Arm board dimension : ≥ (400mm x 140 mm)\n• Capacity(loading): ≥180 kg\n• Power: 230V,50Hz, \n",5]
[11,"Blood bag stripper \nMachined aluminum\nTefl on rollers with reinforced pins\nElectroless Nickel Coating; durable with a chrome-like fi nish\n",5]
[12,"\nPlatelet aggregation  analyzers\nPURPOSE :  Platelet aggregation \nanalyzers are used to evaluate the adhesion of platelets to one another\nTESTS : Aggregation, Leukocyte also adhesion, Ristocetin cofactor , Other tests should be specified\nMETHODOLOGY  : Optical ,impedance or both, flow cytometer\nCHANNELS : Multi channels (4-8)\nDETECTOR AMPLIFIER :  Photodiode\nSPECIMEN : Diluted or undiluted whole blood, PRP, WBC  \nSPECIMEN VOLUME, μL : 250, 450,1000\nPREHEAT INCUBATORS : 8-16\nTEMPERATURE RANGE, °C adjustable : Yes\nTEMPERATURE INDICATOR : Meter  \nSTIRRER  Speed adjustable : yes\nAUTOMATIC CALIBRATION : Yes \nAUTOMATIC SETTING : Yes\nMONITORING OF STIRRER AND TEMPERATURE :Yes\nOUT PUT : Color Monitor (touch screen )and laser jet printer \nSOFTWARE : interface with Windows based system computer \nPOWER REQUIRMENTS : 230V,50Hz, \nADDITIONAL REQUIRMENTS:\n- Temperature range, detector sensitivity and stability should be reported by the manufacturer\n- High degree of precision, with high sensitivity instrument \n- Built-in data management systems that allow quality control (QC) and patient test values\n- Automatic  steps such as reagent addition and dilution to eliminate sample preparation -   steps (e.g., separating PRP from whole blood) to reduce the risk of operator contamination \n- System should be with  sufficient optical sensitivity to distinguish between the PRP and PPP samples\n- User defined templates\n-Electronic pipette :included",1]
[13,"Coagulo meter \nFeatures\n1.Easy operation with LCD display\n2.With cold light source technology, long lifespan \n3.Automatically hinting time for incubation time and adding reagent, automatically detecting reagent adding in reagent cup, and automatically start magnetic mixing device and testing timing. \n4.Built-in magnetic stirrer ensuring the sample reagent mixing evenly in the measure. \n5.The result is not affected by abnormal samples, jaundice、hemolysis、hyperlipemia samples can be measured accurately. \n6.The instrument has two detectors, which can be test items selectable or single-item with double tests in the same interface. \n7.Up to 10000 sample results can be stored.\n8.Built-in computer for many types of conversions (S,%, PTR, TNR, g / L), automatically print out the results\n\nSpecification\nMeasuring pricinple: transmission turbidimetry \nMeasuring channel: double channels \nSample warm-up position: 15\nReagent warm-up position: 2\nSample volume: 50μl~100μl\nReagent volume: 50μl~200μl\nData entry: keyboard \nTemperature controlled: 37℃± 0.3℃ \nDisplay: back-illuminated liquid crystal\nPrinter: built-in micro printer\nPower supply: 230V,50Hz, , ,50W",1]
[14,"Aggrego meter \nMeasuring principle: Photometric, high resolution photometer (740 nm) up to 3 OD\nMeasuring unit: Controlled at 37.4 °C +/- 0.3 °C\n4 independently operating measuring cells for simultaneous determination of different tests\nMeasuring sensitivity down to 30.000 Thrombocytes/μL\n4 reagent positions, one stirred\n2 x 15 cuvette incubations positions\nTest volume: Micro cuvette for single use\nTotal min. required 150 μL max. allowable 300 μL\nDisplay: Touch sensitive display\nInterfaces: 2 x RS 232 C / ChipCARDTM / MMC-Card / LAN\nPrinter: External printer via LAN (optional)\nDimensions/Weight: 25.0 x 32.5 x 14.0cm (9.8x12.8x5.5 inch), 5.3 kg (11.7 lb)\nAssays: ADP, Adrenalin, Collagen, Ristocetin, Arachidonic Acid\nHIT von Willebrand Ristocetin Cofactor assay (optional)\nSoftware features: All reagents individuall editable\nOn board storage capability up to 15 assays\nReal time reaction curve recording\nAutomated calculation of max./min aggregation, shape\nchange, lag phase, desaggregation, area and more\nMax. measuring time up to 3.600 sec. prgrammable\nData transmission to MMC-Card or external PC software",1]
[15,"Sterile connecling device ( tube welder) \nTechnical specifications\nPurpose: to seal blood bags tubes \n-Heavy duty and compatible design \n-Adjustable sealing time with range from 0.5ses to 5sec or less for permanent sealing\n-Power: 80/50 Ohm maximum\n-Frequency: 40.68MHz OR 40.0MHz\n-Electric shock protections \n-4 segments sealing or more  \n-Input Power: 220-230V,50-60Hz \n-Operating temperature: 5-40oC\n-Storage temperature: -5 - 60 oC\n",1]
[16,"Plasma  / low hemoglobin meter \nPrinciple:  Modified azidemethemoglobin reaction dual wavelengths (570 nm and 880 nm)\nfor compensation of turbidity\nSample Material : Plasma, serum and aqueous solutions,or stored erythrocyte suspensions\nMeasurement:  0–30.0 g/L (0–3.0 g/dL,\nRange :  0–3000 mg/dL, 0–1.9 mmol/L)\nResults:  Within 60 seconds\nSample Volume  :  ~20 µL\nDimensions:  160 × 210 × 90 mm\n(6.29 × 8.26 × 3.54 inches)\nWeight:  1kg (2.0 pounds) with batteries installed\nStorage Temp;  Analyzer: 0–50 °C (32–122 °F)\nMicrocuvettes: unopened 15–30 °C\n(59–86 °F); three-month open vial\nstability\nOperating Temp.:  15–30 °C (59–86 °F)\nPower:  AC Adapter or 4 AA batteries\n",1]
[17,"\nWhole blood processing system  \n Application :used to fully automate and standardize whole blood process produce blood cells, plasma and platelet components\n Where used :main blood bank \n Configuration :compact desing .\n System procedures \no Protocol 2 component : process whole blood units ≥4 at (≤25 min ) to protocol (≥4 RBCs units) &(4≥plasma units) & (≥4 residual leukocyte units). \no Protocol 3 component : process whole blood units ≥4 at (≤ 25 min) to protocol (≥4 RBCs units)& (4≥ residual leukocyte units) units platelet.\n Data management : yes\n Integrated processing set : the same set is used for collection and processin .\n Start up _kit : 100set /day for one week.\n Power supply :medical approved power supply board \n Line voltage: 220/240VAC, 50Hz, three pin (G type) plug\n Voltage stabilizer circuit(or device) : yes (170-260VAC to 220/240VAC)\n cUPS : yes for ≥15 minutes back up \n Noise level : 50-65hz\n Environmental requirements : the supplies confirmation the equipment suit for working in the climate conditions of Iraq in terms of temperatature &humi. \n",2]
[18,"\nFull automated viral screening \nMETHODOLOGY: Enzyme-linked\nTESTS performed:\nAntianemics ,Antiarrhythmics, Antiasthmatics, Antibiotics, Anticonvulsants, Antineoplastics\nChemistries, Diabetes , Drugs of abuse , Endocrine function , Ethanol , Thyroid markers\nImmunoglobulins, Immunosuppressants, Nucleic acid, Proteins, Toxoplasma, Tumor markers\nViruses\nSAMPLE TYPE: Serum, plasma, urine, other\n",1]
[19,"\nFull automated blood grouping test\n\nPurpose : Blood grouping systems used for ABO grouping and subgrouping, RH and other red cell phenotyping, and antibody detection.\nThroughput, samples/hr : to user requirements; higher throughput preferred\nAnalysis time : Less  than 60 min\nABO forward/reverse : Yes\nRH : Yes \nABO subgrouping : Opational prefered \nweak D : Opational prefered\nother red cell : Opational prefered \nphenotyping , antibody screens\nSpecimen type : EDTA plasma, serum , other to be specified\nSample volume, µl : smaller sample volume preferred\nReagents volume, µl : smaller reagent volume preferred\nAgglutination temperature : ambient and 37°c\nReaction readings : Yes\nSystem capabilities\nProcessing mode : batch, random access , other to be specified\nstat testing : Optional\nclot detection : Yes\nauto dilution : Yes\nAgitation : Yes\nCentrifugation : Yes\nData management \ncomputer interface : Bidirectional\nIntegrated computer : Optional ( preferred )\nbar-code reader : Yes\nPrinter : Yes\nWater supply : Yes , Self-pump or container\nAlert indicators : Yes\nLine power, VAC : 220/240 VAC, 50/60 HZ\n",1]
[20,"\n\nCool spine centrifuge ( Refrigerated Centrifuge )\nApplications:\nTo separate suspended particles from a liquid or to separate liquids of different densities especially for temperature sensitive materials\n\n - Max rpm: ≥ 10,000\n - Speed adjustment: ≤ 10 steps, microprocessor control \n- Rotors: horizontal fixed angel, preferable to accept different rotors \n- Temperature range, °c :-10 to +40 \n- Temperature accuracy, °c : ≤3\n- Timer range, min: 0 to ≥60\n- Timer settings: timed, continuous\n- Breaking, type: Electric, with time breaks setting \n- Brushes: brushless \n- Safety interlock: lid interlock \n- Imbalance, end of run, over speed alarms \n- Display: Parameters displayed Speed, RCF, time, temperature\n- Self diagnostic facility \n- LINE POWER, 230V,50Hz, \n\nOther specifications:\n- Shut off the motor when the lid is opened or keep the lid latched until the rotor has stopped\n- Inner protective cover over the rotor assembly\n- Should have clear, easily read visual indicators or displays; audible signals or alarms for device malfunction\n- Refrigerants should be CFC-free\n\n",2]
[21," laboratory Balance \nspecifications\nPurpose: For precisely weighting of ultra-low weight materials for laboratory uses \n\n-Digital scale \n-Rang should read up to 1000g\n-Sensitivity: 1ng (0.001g)\n-Working temperature: 10-400C, storing temp 0-50oC\n-Shock resistance and overload protection \n-Summation of weighting results \n-Automatic calibration \n-Capability of Different units weighting \n-Balance pan :≥ 30cm\n-Anticorrosion material and easy to clean \n-Power requirements :230V,50Hz, \n\n",2]
[22,"\nMicro centrifuge\n specifications\n Purpose:  to separate suspended particles from a liquid or to separate liquids of different densities\n\nROTATIONAL SPEED :> 8000 rpm with fixed setting \nNumber of tubes (type)  : 24 (75 mm capillary) \nROTORS :Horizontal \nTIMER : Range, min : 0 to ≥60 with increment ≤ 1min\nBRAKING TIME, sec : 120 \nBRUSHLESS  :Yes \nNOISE LEVEL, dB : <70 \nLID INTERLOCK: Yes \nALERT INDICATORS: Audible imbalance, leakage current, lid open, end of run \nOther specifications :Automatic breaking , air vent on the cover, LED display and a keyboard, tachometer to indicate the speed in rpm should be available , inner safety lids that fit over the rotor or tube covers that contain the sample material should be  available, Gaskets with high quality material for tight seal, Bowl drains, the centrifuge should allow calibration by phototachometer manner with lid closed, should use at least an 8 cm rotor radius, a 50 μL sample, and a 75 mm capillary tube, inner protective cover over the rotor assembly\nEnvironmental requirements: the equipment suitable for working in the climate conditions of Iraq in terms of temperature & humidity\nLINE POWER, VAC : 230V,50Hz, \n\n",2]
[23,"Laboratory Incubator \n-  Purpose: used for growing microorganisms and cells that require oxygen for growth (e.g., bacteriology, virology, cultures) in hospital lab, specialist health center lab. & lab. of  Primary health care unit.\n- Configuration: tabletop \n- Inner chamber stainless-steel and out chamber heavy duty material.\n- Type door :  glass or solid door\n-  Key Door lock :yes \n- Capacity  , L :  150 – 200 L \n- Number of Shelves : ≥  2 , stainless steel\n- Temp is controlled by thermostat.\n- Temperature range , °C   :   5  to 40 °C\n- Temperature Uniformity  +/- 0.5°C\n- Accuracy :  ± 1  °C\n- Illumination : internal light \n- Control system : microprocessor controlled temperature regulation   \n- display : digital\n- Alarms: audio & visual, Temperature, power failure\n- the heated air circulates through the chamber by fans.\n- All the control switches & Pilot lamps are fitted on the front pannel\n- Power requirements: 230V,50Hz, \n- Environmental requirements: the equipment suitable for working in the climate conditions of Iraq in terms of temperature & humidity\n",2]
[24,"\nElecetrecal oven up to 200 C\nTemperature range:      50°C ~ 300°C  \nTemperature control:  Thermostatically controlled\nInside material: Double /Triple walled Construction inner chamber made of Stainless steel with  supports on the three sides for two steel  shelves\nOutside material : Mild steel sheet painted/powder coated Insulated stainless steel door with locking and rear zinc plated steel\nCapacity :  > 50  Itr.\nSafety Alarm : Yes (Audible and visual) - Optional automatically shut down protection circuit                                             \nAir circulating fan        Yes\nDigital timer                   Yes \nPower supply        220/240 VAC, 50Hz \nOver current & over heated protection circuit\n Three Pin UK type Plug  Voltage corrector/stabilizer circuit of (Input  160-260 V and output 220-240 V and 50 Hz)\nOther specification\n on/off rocker switch with indicator \nAdjustable wire mesh nickel plated trays/aluminum trays\nDigital indicator cum controller \nThe gap between the two walls filled with glass wool insulation \n Easy-to-clean interior \n Forced air circulation by quiet air turbine/Fan to ensure uniform  temperature \n Fitted with load indicator and safety thermostat take over      indicator   lamp. LCD/LED Indicator \nOutput available for data acquisition \nThe unit shall be capable of being stored continuously in ambient temperature of 0 ~ 70° C and relative humidity of 15-90%\nThe unit shall be capable of operating continuously in ambient \ntemperature of 0 ~ 50° C and relative humidity of 15-90%\n\n",1]
[25,"\nEliza\nWavelength Range: 400 to 750 nm\nFilters: 10 nm half-bandwidth interference filters.\nUser-accessible filter wheel. Up to 5 filters may be\ninstalled on the instrument at one time.\nFilters supplied: 405 nm, 450 nm, 490 nm and 630 nm.\n• The following specifications apply to 96-well, flat- or round-bottom plates,\nsingle-wavelength measurements with a 50-second read (normal read\nmode).\nAbsorbance Measurement\nRange:\n0.000 to 3.000 OD\nAccuracy: ± 1.0% ± 0.010 OD from 0.000 to 2.000 OD\n@ 405 nm\nLinearity: ± 1.0% from 0.000 to 2.000 OD at 405 nm\n± 3.0% from 2.000 OD to 3.000 OD @ 450 nm\nRepeatability (STD): ± 0.5% ± 0.005 OD from 0.000 to 2.000 OD @ 405 nm\n• The following specifications apply to 96-well, flat- or round-bottom plates,\nsingle-wavelength measurements with a 30-second read (rapid read mode).\nAbsorbance Measurement\nRange:\n0.000 to 3.000 OD\nAccuracy: ± 2.0% ± 0.020 OD from 0.000 to 2.000 OD @ 405 nm\nLinearity: ± 2.0% from 0.000 to 2.000 OD @ 405 nm\nRepeatability (STD): ± 1.0% ± 0.010 OD from 0.000 to 2.000 OD @ 405 nm\n• Light Source: Tungsten gas-filled bulb\n• Dimensions: 42 cm x 38 cm x 18 cm (16.5\" deep x 15\" wide x 7\" tall)\n• Weight: 8 kg (18.5 lb. Maximum)\n• Environment: Operating temperature 18° to 40°C\n• Humidity: 10% to 85% noncondensing\n• Power Supply: Input 100 to 240 V~ ± 10.0% @ 50 to 60 Hz\nOutput +24 VDC, 2.1 A \n",1]
[26,"Shaker \n• Purpose: used in laboratory for gentle rolling motion to provide mixing at steady even rate to prevent sedimentation of samples.\n• Motion : rolling & tilting \n• Speed range : ≥ (10 – 40 rpm )\n• Adjustable speed : yes \n• Number of roller : ≥ 5 \n• Display : LCD \n• Controls : digital & solid state \n• Timer : yes \n• Maximum  load : ≤ 1.6 kg \n• Roller : robustly constructed and designed for easy  cleaning , rubber tough , high–impact rollers \n• Rotor drive : metal with stainless steel drive pins \n• Tube type : accept various tube diameter and length \n• Line power : 230V,50Hz, \n• Environmental requirements : The equipment suitable for working  in the climate conditions of Iraq in terms of temperature & humidity\n",2]
[27,"fully automatic high-speed pre & post vacuum steam sterilizers autoclave  . (25 )Litter                                                                                               \n- chamber is made of stainless steel 316L (conforms to the Pressure Equipment Directive (PED).  Built-in  dry membrane vacuum pump,  Built –in steam generator , Vacuum pump for evacuation of the chamber, Water reservoirs for process and used water retention, , LCD 4.5’’ screen , Water pump for water circulation,Condenser and collector for condensation and collection of water. Supplied with : RO  Reverse-Osmosis system 100G/day , One Stainless Steel Tray Holder.                        Meet the following  : EN 13060:2004+A1:2009+A2:2010 ‐ Small Steam Sterilizers; Pressure Equipment Directives (PED) 97/23/EC; EMC Directive 2004/108/EC Article 7 (1), EN 61326: 1997 + A1(98) + A2(01) + A3(03), minimum immunity requirements, Class A; EN 62366:2008, EN ISO 10993-1:2009; EN 1041:2008, EN ISO 15223-1:2012 , ISO 17665-1:2006; EN ISO 14971:2012 – Risk Management for Medical device; EC ‐ Certificate Number: 784 CE (CE 0473 for medical devices); ISO 13485:2012 - Quality systems - Medical devices -Particular requirements for the application of ISO 9001",1]
[null,"GRAND TOTAL IN US$ "]

{"sheet":"Sheet2","index":1,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[]}

{"sheet":"Sheet3","index":2,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[]}