Anexo 1

SOLICITUD DE COTIZACIÓN (RFQ)

| CENTRO REGIONAL DEL PNUD PARA AMÉRICA LATINA Y EL CARIBE | FECHA: Marzo 25, 2020 |
| --- | --- |
|  | REFERENCIA: 17230 RSC 2020 |

Especificaciones técnicas

| Artículo | Especificación | Empaque requerido | Favor indicar Stock disponible al momento de ofertar*** |
| --- | --- | --- | --- |
| MASK - SURGICAL MEDICAL/SURGICAL MASK 3 PLY | Medical/surgical mask, good breathability, internal and external faces should be clearly identified, structured design that does not collapse against the mouth (e.g. duckbill, cup-shaped) 3ply disposable face mask | 2000 pcs / carton 40 x 50 pcs/ pack |  |
| MASK - SURGICAL MEDICAL/SURGICAL MASK TYPE IIR | Medical/surgical mask, high fluid resistance, good breathability, internal and external faces should be clearly identified, structured design that does not collapse against the mouth (e.g. duckbill, cup-shaped) EN 14683 Type IIR performance ASTM F2100 level 2 or level 3 or equivalent; • Fluid resistance at minimum 120 mmHg pressure based on ASTM F1862-07, ISO 22609, or equivalent • Breathability: MIL–M-36945C, EN 14683 annex C, or equivalent Filtration efficiency: ASTM F2101, EN14683 annex B, or equivalent | und |  |
| MASK - RESPIRATOR PARTICULATE RESPIRATOR FFP2 / N95 SIZE S | Good breathability with design that does not collapse against the mouth (e.g. duckbill, cup shaped) "N95" respirator according to US NIOSH, or "FFP2" according to EN 149, Size S | und |  |
| MASK - RESPIRATOR PARTICULATE RESPIRATOR FFP2 / N95 SIZE M | Good breathability with design that does not collapse against the mouth (e.g. duckbill, cup shaped) "N95" respirator according to US NIOSH, or "FFP2" according to EN 149; Size M | und |  |
| MASK - RESPIRATOR PARTICULATE RESPIRATOR FFP2 / N95 SIZE L | Good breathability with design that does not collapse against the mouth (e.g. duckbill, cup shaped) "N95" respirator according to US NIOSH, or "FFP2" according to EN 149, Size L | und |  |
| MEDICAL / SURGICAL MASK TYPE IIR (EL SALVADOR LOT) | DOUBLE LINING RECTANGULAR SURGICAL MASK, FITTED, WITH NASAL SUPPORT, DISPOSABLE • Rectangular double lined surgical mask • With elastic bras • Three-fold • With moldable bar, anatomically adaptable to the nose on the upper edge of the mask (nasal support) • Disposable Warranty: 3 months | und. |  |
| MASK - RESPIRATOR PARTICULATE RESPIRATOR FFP2 / N95 (EL SALVADOR LOT) | "HIGH EFFICIENCY FILTER (MASK) WITH NIOSH N95 CERTIFICATE OR HIGHER, WITH NASAL ADAPTER, TWO ELASTIC CLAMPS, FACIAL SEALED ADAPTATION, WITHOUT EXHALATION VALVE • High filtering efficiency respirator (mask) with NIOSH N95 or higher certificate • With nasal adapter • With two elastic bras • No latex content, no staples for greater comfort and safety. • Soft sponge nose pad with airtight fitting • Without exhalation valve • Different sizes or measurements (S, M, L or XL) Warranty: 12 months. " | und. |  |
| HAND SANITIZER 100 ML | Alcohol-based hand rub | 1 botella de 100ml |  |
| HAND SANITIZER 500 ML | Alcohol-based hand rub | 1 botella dispensadora de 500ml |  |
| GLOVES S (SMALL) | GLOVE, EXAMINATION, extended cuff, nitrile, powder-free, non-sterile single use, SMALL Cuff length preferably reach mid-forearm (eg. minimum 280mm total length) reaching well above the wrist • EU standard directive 93/42/EEC Class I, EN 455, • EU standard directive 89/686/EEC Category III, EN 374, • ANSI/ISEA 105-2011, • ASTM D6319-10 • or equivalent | Caja de 100 |  |
| GLOVES M (MEDIUM) | GLOVE, EXAMINATION, extended cuff, single use, nitrile, powder-free, non-sterile, MEDIUM Cuff length preferably reach mid-forearm (eg. minimum 280mm total length), reaching well above the wrist • EU standard directive 93/42/EEC Class I, EN 455, • EU standard directive 89/686/EEC Category III, EN 374, • ANSI/ISEA 105-2011, • ASTM D6319-10 • or equivalent | Caja de 100 |  |
| GLOVES L (LARGE) | GLOVE, EXAMINATION, extended cuff, single use, nitrile, powder-free, non-sterile, LARGE Gloves, examination, non-sterile. Cuff length preferably reach mid-forearm (eg. minimum 280mm total length), reaching well above the wrist • EU standard directive 93/42/EEC Class I, EN 455, • EU standard directive 89/686/EEC Category III, EN 374, • ANSI/ISEA 105-2011, • ASTM D6319-10 • or equivalent | Caja de 100 |  |
| DIGITAL THERMOMETER | Clinical thermometer Digital thermometer Celsius scale. Safe to use, no glass, no mercury. Measurement range: 32°C to 43°C. Accurate measurement: +/- 0.1°C between 35°C to 41°C. Liquid crystal display, easy to read. Beep sound and switch off. Waterproof for ease of cleaning. Battery powered (min. 1000 measurement). Low battery indicator. Supplied with battery. Supplied with clear instructions for use/preventive maintenance. Conforms to EN 12470-3 | und |  |
| NON-CONTACT INFRARED THERMOMETERS (NCITS) | Infrared thermometer Displays patient temperature by measurement of infrared radiation from the skin Device must be reusable, with sterilizable surface Display should be easily readable in all levels of ambient light Specified accuracy to be better than 0.3 deg C * Measurement range at least to 25 at 42 deg C * * High / low patient temperature display feature preferred * Auto power off required after minimum of 1 minute * ‘Out of range’ indication required * Response time to steady reading < 5 seconds required Specified accuracy to be better than 0.3 deg C Display graded in maximum 0.1 deg C steps. High/low patient temperature, low battery, malfunction, °F or °C. | und |  |
| NON-CONTACT INFRARED THERMOMETERS (NCITS) (EL SALVADOR LOT) | INFRARED THERMOMETER (NO CONTACT • Infrared measurement technology, taken from the patient's forehead, without contact. • Temperature result in less than 10 seconds. •Date and time indicator. • Celsius and Fahrenheit Degrees indicator. • Acoustic signal. • Useful range: 34 ° C to 42 ºC. • Fever alarm. • Measurement accuracy on the forehead: ± 0.2ºC • Memory for at least 50 measurements. • Backlit LCD display, 3 digits. •Auto power off. • With battery replacement indicator. Included Accessory: Case Electrical Characteristics: • Voltage: any of the following options:  By rechargeable batteries, with battery charger at (120 ± 10%) VCA, Frequency: 60 Hertz, single phase.  Pill type include an additional one. Mechanical Characteristics: • Strong and easy to clean casing. • Weight: no more than 100 grams. Standards and Regulations • Manufactured under quality management standards: ISO 13485 or equivalent. (Present supporting certificate). • Approved for marketing by one of the following organizations: JPAL, PMDA, JIS, (Japan), Directive 93/42 EEC of the European Community or FDA (USA). (Submit supporting documentation). Technical support • The company must have a biomedical technical service department to guarantee qualified technical support for the equipment offered. • MINSAL reserves the right to verify the information received in this regard. | und. |  |
| GOWN S (SMALL) | Single use, fluid resistant, disposable, length mid-calf to cover the top of the boots, light colors preferable to better detect possible contamination, thumb/finger loops or elastic cuff to anchor sleeves in place. • Option 1: fluid penetration resistant: EN 13795 high performance, or AAMI PB70 level 3 performance or above, or equivalent • Option 2: blood borne pathogens penetration resistant: AAMI PB70 level 4 performance, or (EN 14126-B) and partial body protection ( EN 13034 or EN 14605), or equivalent | und |  |
| GOWN M (MEDIUM) | Single use, fluid resistant, disposable, length mid-calf to cover the top of the boots, light colors preferable to better detect possible contamination, thumb/finger loops or elastic cuff to anchor sleeves in place. • Option 1: fluid penetration resistant: EN 13795 high performance, or AAMI PB70 level 3 performance or above, or equivalent • Option 2: blood borne pathogens penetration resistant: AAMI PB70 level 4 performance, or (EN 14126-B) and partial body protection ( EN 13034 or EN 14605), or equivalent | und. |  |
| GOWN L (LARGE) | Single use, fluid resistant, disposable, length mid-calf to cover the top of the boots, light colors preferable to better detect possible contamination, thumb/finger loops or elastic cuff to anchor sleeves in place. • Option 1: fluid penetration resistant: EN 13795 high performance, or AAMI PB70 level 3 performance or above, or equivalent • Option 2: blood borne pathogens penetration resistant: AAMI PB70 level 4 performance, or (EN 14126-B) and partial body protection ( EN 13034 or EN 14605), or equivalent | und. |  |
| FULL FACESHIELD | Made of clear plastic and provides good visibility to both the wearer and the patient, Adjustable band to attach firmly around the head and fit snuggly against the forehead, Fog resistant (preferable), Completely cover the sides and length of the face, May be re-usable (made of robust material which can be cleaned and disinfected) or disposable. EU standard directive 86/686/EEC, EN 166/2002, ANSI/ISEA Z87.1-2010, or equivalent | und. |  |
| GOOGLES | Good seal with the skin of the face, Flexible PVC frame to easily fit with all face contours with even pressure, Enclose eyes and the surrounding areas, Accommodate wearers with prescription glasses, Clear plastic lens with fog and scratch resistant treatments, Adjustable band to secure firmly so as not to become loose during clinical activity, Indirect venting to avoid fogging, May be re-usable (provided appropriate arrangements for decontamination are in place) or disposable. EU standard directive 86/686/EEC, EN 166/2002, • ANSI/ISEA Z87.1-2010, or equivalent | und. |  |
| MASK, SURGICAL – PATIENT | Surgical mask, good breathability; internal and external faces should be clearly identified Type I. EN 14683 any type including Type I • ASTM F2100 minimum level 1 or equivalent | Und. |  |
| COVERALL WITH HOODS | Mask, medical, good breathability, internal and external faces clearly identified * EN 14683 any type including Type I * ASTM F2100 any level | Und. |  |
| APRON, HEAVY DUTY | Straight apron with bib, Fabric: 100% polyester with PVC coating, or 100% PVC, or 100% rubber, or other fluid-resistant coated material. Waterproof, sewn strap for neck and back fastening Minimum weight: 300 g/m2 Covering size: 7090 cm (width) x 120–150 cm (height) Reusable (provided appropriate arrangements for decontamination are in place) EN ISO 13688 EN 14126-B and partial protection (EN 13034 or EN 14605) • EN 343 for water and breathability or equivalent | Und. |  |
| GOWN (HAITI) | Single-use, length mid-calf EU PPE Regulation 2016/425 and EU MDD Directive 93/42/EEC FDA Class I or II medical device, or equivalent EN 13795 any performance level, or AAMI PB70 all levels acceptable, or equivalent | Und. |  |
| GLOVES, CLEANING | Outer glove should have long cuffs, reaching well above the wrist, ideally to mid-forearm, Minimum 280 mm total length Sizes: small, medium, large Reusable Puncture-resistant, FDA compliant | Und. |  |
| PATIENT VENTILATOR, FOR CRITICAL CARE | Tidal volume up to 1000 Ml • Pressure (inspiratory) up to 80 cm H2O • Volume (inspiratory) up to 120 L/min • Respiratory rate: up to 60 breaths per minute • Synchronized intermittent mandatory ventilation (SIMV) respiratory rate: up to 40 breaths per minute. • CPAP/PEEP up to 20 cm H2O • Pressure support up to 45 cm H2O • FiO2 between 21% and 100% • Inspiratory and expiratory times up to at least 2 sec and 8 sec respectively • I:E ratio from 1:1 to 1:3 Modes of ventilation: • Volume controlled • Pressure controlled • Pressure support • SIMV with pressure support • Assist/control mode • CPAP/PEEP Alarms are required: FiO2, minute volume, pressure, PEEP, apnea, occlusion, high respiration rate, disconnection. System alarms required: power failure, gas disconnection, low battery, vent inoperative, self-diagnostics. If an alarm silencing feature is incorporated, it must be temporary and clearly displayed when activated. Air and externally supplied oxygen mixture ratios fully controllable; inlet gas supply (O2) pressure range 35 psi to 65 psi; medical air compressor integral to unit, with inlet filter ISO 80601-2-80 and ISO 80601-2-79 or equivalent | und. |  |

Unidad de Adquisiciones PNUD-RBLAC

Marzo 26, 2020