{"text_format":"2","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":50,"n_cols":5,"rows_kept":50,"cut":false,"hidden":false,"merged":[[3,2,3,5],[16,2,16,5],[42,2,42,5],[48,2,48,5]]}
[]
[null,"Art no","Article","MODEL ","Price "]
[]
[null,null,"Equipment"]
[null,1,"Microbiology Safety Cabinet, Clas II Type A2 ","Technical specification of Microbiology Safety Cabinet, Class II Type A2\nA ventilated containment device with:\nWorking zone made of stainless steel and surrounded by negative pressure. \nMain body from steel with anti-bacteria powder coating. \nFront window: two-layer laminated toughened glass>5mm, anti UV\nAirflow pattern: approximately 70% of air is recycled and approx 30% exhausted\nExhaust system type: recirculating — filtered exhaust into room\nInflow velocity: ≥ 100 FPM (NSF 49),  ≥ 0.40 m/s (EN12469)\nDownflow velocity: Not defined (NSF 49); 0.25–0.50 m/s (EN12469)\nMotorized front window. \nHEPA Filter:  two, 99.999% efficient at 0.3 um. Filter life indicator.\nUV lamp with UV timer, UV life indicator.\nA daylight lamp for illumination of the working area \nAudio and visual alarm (filter replacement, window over height, abnormal air flow velocity etc).\nLCD display: exhaust filter and downflow filter pressure, filter and UV lamp working time, inflow and downflow velocity, filter life, humidity and temperature, system working time etc.\nExternal Size (W*D*H): 1100*750*2250mm\nDelivery, installation, and commissioning of equipment.\nTraining of use the equipment is included \nWarranty service for at least 12 months.\nTechnical support available\n"]
[null,2," PCR Master Mix Cabinet  with build in UV light ( PCR workstation) ","Technical specification of the PCR cabinet (DNA/RNA UV-cleaner box)\nBenchtop cabinet designed for clean operations with DNA samples protected against contamination. \nMade of metal framework and working surface of stainless steel.  \nEquipped with an open UV lamp installed in the upper hood  \nDigital time setting of direct UV exposure: 1 min–24 hrs / non–stop (increment 1 min)\nA daylight lamp for illumination of the working area.\nUV-recirculator: 1x25W (efficiency >99% per 1 hour)\nAutomatic switch- off of UV-lamps when the protective screen is opened\nOptical transmission: 95%\nLow noise, low energy consumption\nNominal operating voltage: 100-240 V, 50/60Hz\nOverall dimensions (W×D×H): 700 × 580 × 555 mm\nDelivery, installation, and commissioning of equipment.\nTraining of use the equipment is included \nWarranty service for at least 12 months.\nTechnical support available"]
[null,3,"Laboratory refrigerator ","Technical specification for laboratory refrigerator\n\nTemperature Range:    2°C to 9°C\nCapacity (Metric):  minimum 156L\nCertifications/Compliance: CE\nDefrost: Automatic\nGreen Features: Energy efficient\nElectrical Requirements: 100 - 240 V 50-60 Hz\nWarranty service for at least 12 months.\nDelivery, installation, and commissioning of equipment.\n\n"]
[null,4,"Laboratory freezer , ","Technical specification for laboratory ultra low temperature freezer\n\nAt least minus 70 °C \nTemperature monitoring\nTemperature Control Range -50°/-86°C \nDisplay \nUpright \nCapacity – at least 90L\nAlarms available:\n-Visual and acoustic alarms\n-Adjustable high/low temperature alarm \n-Power failure alarm \n-Open door alarm\n-Probe failure alarm\nKey lock \nFully equipped with racks and boxes\nCO2 backup system 50Hz\nWarranty service for at least 12 months.\nDelivery, installation, and commissioning of equipment."]
[null,5,"Shaker Vortex","Technical specification of Vortex Mixer\nLaboratory benchtop vortex/mixer \n Wide speed range of 0-3000 rpm\nOrbital Diameter 6mm\nAvailable:  touch mode and continuous mode \nOperating Frequency: 50/60Hz\nWarranty service for at least 12 months.\nDelivery, installation, and commissioning of equipment.\nTraining of use the equipment is included\nTechnical support available"]
[null,6,"Dry Block Termostat","Technical specification of Dry block thermostat\nDry block thermostat designed for long incubation processes at various temperatures. \nWith:\nUniversal aluminium block  who can accommodate  at least one (1.5 ml microtubes)  or 3 types of microtubes' (  1.5 ml + 0.5 ml +0.2 ml microtubes)\nSimultaneous indication of set and actual temperature and time\nTemperature setting range: +25°C ... +100°C\nTemperature stability:  ±0.1°C\nTemperature setting resolution: 0.1°C\nDigital time setting: 1 min–96 hrs / non–stop (increment 1 min)\nDisplay: LCD\nWith timer sound signal\nNominal operating voltage: 230 V, 50/60 Hz or 120 V, 50/60 Hz\nWarranty service for at least 12 months.\nDelivery, installation, and commissioning of equipment.\nTraining of use the equipment is included\nTechnical support available"]
[null,7,"Microcentrifuge; Refrigerated","Technical specification of combined Minicentrifuge – vortex\n\nEquipment for simultaneous mixing and separation of samples, using centrifuge and mixing modules, located on the common spin–module. \nWith or without lid. \nMaximum RCF: from 500 to 800 × g. \nWarranty service for at least 12 months.\nDelivery, installation, and commissioning of equipment.\nTraining of use the equipment is included\nTechnical support available"]
[null,8,"Minicentrifuge ","Technical specification of minicentrifuge\nMinicentrifuge is a benchtop centrifuge for quick spin downs of microtubes and PCR tubes. \nEight position microtube rotor for 1.5/2.0 mL tubes\nWhen the lid is closed, the rotor instantly reaches 6,000 rpm (2000 × g)\nThe rotor quickly decelerates to a stop when open the lid. \nWarranty service for at least 12 months.\nDelivery, installation, and commissioning of equipment.\nTraining of use the equipment is included\nTechnical support available\n"]
[null,9,"Autoclave","Technical specification for autoclave\nLaboratory vertical autoclave with top loading intended for steam sterilization of metal instruments, textiles, rubber, plastics, solutions in open and/or semi closed bottles, glass, packed items, spongy materials and laboratory waste. With:\nBuilt in steam generator (with automatic self-cleaning system)\nChamber made of stainless steel with volume from 85 -110 L\nStainless steel steam passivated pipes, stainless steel frame and steam valves \nPower: 400 V; 50/60Hz;\nControl system with multi-color 70x50 mm display, with up to 30 cycle programs. \nTemperature range 105°C to 138°C \nSterilization time range 1 – 99 min \nTwo real-time clocks for cross-checking timing accuracy\nTemperature and pressure regulation and monitoring\nTemperature sensors in the chamber to prevent over-boiling of liquids and explosions of bottles \nCycle control through absolute pressure sensors (control and recording sensor according to EN 285:2015)\nLocking system: automatic pressure door locking system\nDoor lock safety switch\nPressure and Temperature dependent door lock/unlock\nIdentification codes and passwords for access level control\nExhaust steam cooling\nMineral-free water filling inlet\nInternal water reservoir\nWater level indicator\nWater level safety device\nFiltration of the exhaust air through a 0.2 μm biological filter to prevent contamination of the laboratory\nHigh-output vacuum pump (up to 7 kPa of absolute pressure for proper drying)\nFunction of multiple vacuum pumping (99% vacuum)\nAll pipes and piping connection are passivated \nTwo baskets included one regular, one for container with vent holes\nValidation port according to EN 285:2015\nTemperature cut off switch\nOverpressure safety valve\nWheels for easy mobility\nBuilt-in buzzer for audible notification\nSystem of self-diagnostics with alarm indication. Alarm list with description must be included.\nEasily generate PDF reports\nIncludes USB and ethernet connection communication ports. \nSoftware for storing protocols, alarms included. \nMultilevel access to menu (user, admin, service) \nMulti language compatible.\nBuilt-in printer for printing data\nPreset and editable (customizable) sterilization programs.\nAll decontamination sterilizers have to be optionally equipped with special sterilization program for decontamination (for reliable inactivation of viruses and risky or dangerous materials)  \nPreset test programs for Vacuum test and Bowie Dick test according to EN 285:2015.\nFlash memory for up to 100 cycles\nAvailable validation documents on request\nDelivery, installation, and commissioning of equipment.\nTraining of use the equipment is included\nWarranty at least 2 years, preferably up to 5 years\nPost warrant service must be available\nTechnical support available\nMeets the following standards:\n-ISO 13485:2016 Medical devices — Quality management systems — Requirements for regulatory purposes (for the manufacturer of the equipment)\n-EN 285:2015 Sterilization. Steam sterilizers. Large sterilizers\n-ISO 17665-1:2014 Sterilization of Health Care Products -- Moist Heat -- Part 1: Requirements For The Development, Validation And Routine Control Of A Sterilization Process For Medical Devices\n-EN 61010-1: Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 1: General requirements\n-EN 61010-2-041: Safety requirements for electrical equipment for measurement, control, and laboratory Use Part 2-041: Particular requirements for autoclaves using steam for the treatment of medical materials, and for laboratory processes\n-DIRECTIVE 93/42/EEC on medical devices\n-The Pressure Equipment Directive (PED) 2014/68/EU\n-Conforms to PED 97/23 EEC (Europe)\n-Conforms to standards: ASME, UL (USA)\n-Company system compliance with ISO 9001:2008 and ISO 13485:2003\n"]
[null,10,"Computer ","Notebook, i5,GEN 9, 15.6\", RAM 8GB, SSD 256GB, "]
[null,11,"Printer","LaserJet, b/w"]
[null,12,"UPS","back-up UPS, 650 VA"]
[null,null,"Supplies"]
[null,13,"96-Well Reaction Plate, 0.1 mL; for ABI 7500 fast (10 plates/pack)","Pack"]
[null,14,"PCR plate seal (100 pcs/pack)","Pack"]
[null,15,"Sterile aerosol barrier tips 10 µl ","PCR clean for molecular testing (1x96)"]
[null,16,"Sterile aerosol barrier tips 20 µl ","PCR clean for molecular testing (1x96)"]
[null,17,"Sterile aerosol barrier tips 200 µl ","PCR clean for molecular testing  (1x96)"]
[null,18,"Sterile aerosol barrier tips 1000 µl ","PCR clean for molecular testing (1x96)"]
[null,19,"50-200µl adjustable pipettes","50-200µl adjustable pipettes, certificate of calibration by the producer"]
[null,20,"10-50µl adjustable pipettes","10-50µl adjustable pipettes, certificate of calibration by the producer"]
[null,21,"2-20µl adjustable pipettes","2-20µl adjustable pipettes,  certificate of calibration by the producer"]
[null,22,"100-1000µl adjustable pipettes","100-1000µl adjustable pipettes,  certificate of calibration by the producer"]
[null,23,"0.5-10 µl  adjustable pipettes","0.5-10 µl adjustable pipettes,  certificate of calibration by the producer"]
[null,24,"Waste bag stand, 120-130x250-260 mm, steel with ruber feet"," 120-130x250-260 mm, steel with ruber feet"]
[null,25,"Waste bags for waste bag stand","200-220x300-330 mm; 100x1"]
[null,26,"Viral transport medium with flocked swab for collection and transport of clinical specimens containing viruses, including COVID-19; ","At least 1-3 mL UTM in tube; package should includes bags for samples transporting; double pocket;  transparent (approximate dim. 140-160x220-230 mm); Polyethylene bag; with “Biohazard” print and logo (all in one package of 50-200 pcs)"]
[null,27,"Tube racks with handle ","Tube diameter 16 mm, 50 tubes capacity, high density polypropylene"]
[null,28,"Autoclavable Biohazard Bags","Orange bags, 90-95 x 110-120 cm (+/- 5%). High-integrity polypropylene bags withstand 135°C; Pack fo 100 pcs; "]
[null,29,"Autoclavable Biohazard Bags, ","Yellow bags, 90-95 x 120-125 cm (+/- 5%). High-integrity polypropylene bags withstand 135°C; Pack of 100 pcs;"]
[null,30,"CENTRIFUGE Tube, PP 1.5ml, non ster., PCR Clean flat cap, pack-500  ","Microtubes, 1.5ml conical with installed O-ring screw caps, sterilized; 500 pcs in pack;"]
[null,31,"Timer, electronical,  ","LCD display, with stand and alarm; "]
[null,32,"UV lamps for the rooms with remote control","pcs"]
[null,33,"UV lamp mobile","pcs"]
[null,34,"PCR  workup racks","pcs"]
[null,35,"freez (temperature-dependent color-change)  PCR racks","pcs"]
[null,36,"1.5 mL tube racks","pcs"]
[null,37,"Biohazard waste container for infectious waste at least 10L","pcs"]
[null,null,"Reagents"]
[null,38,"FG,FAST 7500 SPECTRAL CAL KIT EACH","pcs"]
[null,39,"FAST 7500 RED SPECTRAL DYE KIT EACH","pcs"]
[null,40,"QIAamp Viral RNA Mini Kit (250)","pcs"]
[null,41,"RT PCR tests for SARS-CoV-19, applicable on ABI 7500 FAST a kit should detect at least 3 genes of SARS-CoV-2 ","TaqPath COVID-19 CE-IVD RT-PCR Kit CE, IVD Mark (1000rxn  per kit)"]
[null,42,"Automated Viral RNA Extraction Kit applicable on ExiPrep™ 48 Dx Bioneer System"," ExiPrep™ Viral DNA/RNA Kit, CE Mark   "]
[null,null,"Personal protective equipment"]
[null,43,"DNA/RNA Decontamination Solution","250 ml pack; eady-to-use solution for eliminating DNA, RNA from surface; non-alkaline and non-carcinogenic agent; "]