{"text_format":"2","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":55,"n_cols":5,"rows_kept":55,"cut":false,"hidden":false,"merged":[[49,3,49,4],[49,0,49,2],[0,0,0,4],[13,0,14,2],[13,3,14,4],[18,0,18,4]]}
["Supply and Delivery of Medical Equipment for  Ramadi Sub PHC Anbaar Governorate                               ( Iraq Al Sijarih, Al Samide, Zawih Satih, Al Sufua PHCs)"]
["No.","Detailed Description","Quantity of packs/kits offered "," Unit Price USD"," Total Price\nUSD"]
[1," Laboratory light Microscope\nbinocular microscope with ball-bearing 5 position objective revolver,the low position coaxial coarse and fine focusing system head ,binocular eyepiece,WF10x objectives 4x 10x  40x   100x oil,mechanical stage withceramic-coatedstage surface and specimen holder and selectable stage driveright or left,<Objectives Magnification:Tension adjustment  :  Coaxial Achromatic, plan achromatic, phase contrast, condenser for bright field dark field and phase contrast ph 1,2,3 ,N.A. 1.25 illuminator : 6 V, 30 W halogen lamp (light blue without heat) with replace- able socket, external power supply . ",8]
[2," Centrifuge, bench-top\nCentrifuge, withelectronic lid lock, swing-out  rotor , imbalance sensors and a rotor recognition program that automatically limits each rotor to its maximum rated speed.  Minimum 24 tubes holders   0 - 4000 RPM , Spins 3ml-15ml test tubes at 4000 rpm , Variable speed and time settings , Brushless motor , Angled and (swinging )tube holder as alternative",4]
[3," Spectrophotometer, \n Visible spectrophotometer,optical bandwith 8 nm , wave length range ( 320nm-1000 nm) , wave length display large seven segment bright LED , absorbance zero setting  0 A & 100% T automatic by press buttom optical system hologiraphic gratings coated optics , light sorces halogen lamp.",4]
[3.1,"Provide on-site installation and training for the  Spectrophotometer "]
[4,"Bilirobinometers\nApplication: To measure the concentration of bilirubin in the blood\nHeavy duty and compatiable design \nPrinter: yes built in \nPower regulator: yes \nConfiguration: yes (Benchtop)\nCapability:\nMethod: Direct spectrophotometery, . ..etc.\nBilirubin measure: Total, ... etc.\nSample type: undiluted serum plasma in centrifuged capilliary tube.\nSample volume: ul: not 30 ul serum or plasma.\nMeasurment time: sec,  ≤ 5 \nCalibration: yes automatic and manual\nFrequancy: for each measurement \nUnits:umol/l or mr/dl\nDisplay: LCD\nComputer interface: yes RS 232\nPowersupply: VAC: 220-240 VAC, 50-60 Hz.\n\n\n",4]
[5,"Water destillation appliance\nWater distillation4Lts/hr  distillation speed , made of stainless steel and two stainless  steel( tanks and condencer) ",4]
[6,"Refrigerator for vaccine store\nTemperature Range +2 / +15 °C\nCapacity more than 300 lt. Door Nr 1 glass door with triple glass panels, Suggested internal fittings nr.4 shelves. Controller ECT-F TOUCH, Alarms Acoustic and Visual for temperature deviations and failures, with automatic recording, Data and Event recording 10 years on SD card; last year on Flash-memory\nTemperature Graph A customizable real time graphic always visible from the home page.\nConnectivity USB-SD-SIM port with frontal access ; Dry Contact; Serial ports RS485 and RJ45\nRefrigeration Type Forced-air, ventilated\nDefrost Automatic with automatic evaporation of the condensate water.",4]
[7," Water bath\nSupply and install WATER BATH with 27 Liter Capacity Independent  high-limit thermostat provides over-temperature protection. double walled metallic construction ,exterior is stainless steel,18/10 stainless steel tub. Electric heating throat resistance covered  in  stainless steel AISI 316 regulation of temp.by means of digital thermostat PID self adjusting ,2800 w power.        Setting temp.range 10°c to 95°c with additional boilig  stage until 100°c , working temp. range from 5°cabove ambient to nominal temp. (maximum temp ), over heat safety device DIN 12877 thermal classes 1 standard.class2 special  equipment ,integral timer for digital programme time selection, the temp. is measured using a Pt 100 sensor with accuracy of(0.1°c for set temp. and ±0.1°c for temp. fluctuation ), the heater is installed outside to avoid the problems due to dirt lime residue.",4]
[8,"Hematocrit Centrifuge\n Hematocrit centrifuge PCV centrifuge) for 24 capillary , digital control  regulared speed  12000 rpm max. with fixed setting  , 230V/50Hz ,24 (75 mm capillary) ,0 to ≥60 with increment ≤ 1min ,BRUSHLESS motor  ,LID INTERLOCK, :Automatic breaking , air vent on the cover, LED display and a keyboard, tachometer to indicate the speed in rpm should be available , inner safety lids that fit over the rotor or tube covers that contain the sample material should be  available, Gaskets with high quality material for tight seal, Bowl drains, the centrifuge should allow calibration by phototachometer manner with lid closed, should use at least an 8 cm rotor radius, a 50 μL sample, and a 75 mm capillary tube, inner protective cover over the rotor assembly. The cost includes providing reagents for 10,000 tests \n\n",4]
[8.1,"On-site installation and training for Hematocrit Centrifuge "]
[9,"Shaker\n Mixer –rotating wave shaker to be used with bottles or closed containers. 5 rollers gives 4 usable places. 25 rpm shaking speed ",4]
["Total :"]
[]
[]
[]
[]
["                                   Supply and Delivery of Medical Equipment for Ramadi Main  PHCs Anbaar Governorate                                                  ( Iraq Al Qtanh, Al Tall, Al Andalus, Al Maleab, Al Warrar, Albu Hussein Al Ali, Al Sura, Husibuh Al Sharqih PHCs)"]
["No.","Detailed Description","Quantity of packs/kits offered "," Unit Price USD","Total Price\nUSD"]
[1," Laboratory light Microscope\nbinocular microscope with ball-bearing 5 position objective revolver,the low position coaxial coarse and fine focusing system head ,binocular eyepiece,WF10x objectives 4x 10x  40x   100x oil,mechanical stage withceramic-coatedstage surface and specimen holder and selectable stage driveright or left,<Objectives Magnification:Tension adjustment  :  Coaxial Achromatic, plan achromatic, phase contrast, condenser for bright field dark field and phase contrast ph 1,2,3 ,N.A. 1.25 illuminator : 6 V, 30 W halogen lamp (light blue without heat) with replace- able socket, external power supply . ",16]
[2," Centrifuge, bench-top\nCentrifuge, withelectronic lid lock, swing-out  rotor , imbalance sensors and a rotor recognition program that automatically limits each rotor to its maximum rated speed.  Minimum 24 tubes holders   0 - 4000 RPM , Spins 3ml-15ml test tubes at 4000 rpm , Variable speed and time settings , Brushless motor , Angled and (swinging )tube holder as alternative",8]
[3," Spectrophotometer, \n Visible spectrophotometer,optical bandwith 8 nm , wave length range ( 320nm-1000 nm) , wave length display large seven segment bright LED , absorbance zero setting  0 A & 100% T automatic by press buttom optical system hologiraphic gratings coated optics , light sorces halogen lamp.",8]
[3.1,"Provide on-site installation and training for the  Spectrophotometer "]
[4,"Automatic ESR analyser\nautomatic ESR analyser with Multi channels, 30 or 60 min program. Tubes with or without vacuum. Tube identification automatically. Sample position detection to minimize user mistakes. RS232 interface connection. The cost includes providing reagents for 10,000 tests.",8]
[4.0999999999999996,"On-site installation and training for automated ESR analyzer\n\n\n"]
[5,"Hematology Analyzer \nParameters\nWBC, LYM, MID, GRA, LYM%, MID%, GRA%, \nHGB, RBC, HCT, MCV, RDWcv, RDWsd, MCH, MCHC, \nPLT, MPV, PCT, PDW.\nTechonolgy Volumetric impadance method, microfluidics\nVial modes:   Closed vial and Open vial mode.   Throughput\n  Closed vial mode 60 tests per hour\n  Open vial mode 50 tests per hour\nPrecision \n(CV, max) WBC: 3%, RBC: 2%, MCV: 2%, HGB: 2%, PLT: 5%\nQC management\nLevey-Jennings chart\nRadar chart\nStorage capacity more than 10.000 test results including histograms.\nUser interface; wide LCD with touch screen\nConnectivity; USB ports, Ethernet, LIS (HL7). The cost includes providing reagents for 10,000 tests.",8]
[5.0999999999999996,"Provide on-site installation and training for the Hematology Analyzer."]
[6,"Hematocrit Centrifuge\n Hematocrit centrifuge PCV centrifuge) for 24 capillary , digital control  regulared speed  12000 rpm max. with fixed setting  , 230V/50Hz ,24 (75 mm capillary) ,0 to ≥60 with increment ≤ 1min ,BRUSHLESS motor  ,LID INTERLOCK, :Automatic breaking , air vent on the cover, LED display and a keyboard, tachometer to indicate the speed in rpm should be available , inner safety lids that fit over the rotor or tube covers that contain the sample material should be  available, Gaskets with high quality material for tight seal, Bowl drains, the centrifuge should allow calibration by phototachometer manner with lid closed, should use at least an 8 cm rotor radius, a 50 μL sample, and a 75 mm capillary tube, inner protective cover over the rotor assembly. The cost includes providing reagents for 10,000 tests \n\n",8]
[6.1,"On-site installation and training for Hematocrit Centrifuge "]
[7,"Refrigerator for vaccine store\nTemperature Range +2 / +15 °C\nCapacity more than 300 lt. Door Nr 1 glass door with triple glass panels, Suggested internal fittings nr.4 shelves. Controller ECT-F TOUCH, Alarms Acoustic and Visual for temperature deviations and failures, with automatic recording, Data and Event recording 10 years on SD card; last year on Flash-memory\nTemperature Graph A customizable real time graphic always visible from the home page.\nConnectivity USB-SD-SIM port with frontal access ; Dry Contact; Serial ports RS485 and RJ45\nRefrigeration Type Forced-air, ventilated\nDefrost Automatic with automatic evaporation of the condensate water.",8]
[8,"Shaker\n Mixer –rotating wave shaker to be used with bottles or closed containers. 5 rollers gives 4 usable places. 25 rpm shaking speed ",8]
[9,"Water destillation appliance\nWater distillation4Lts/hr  distillation speed , made of stainless steel and two stainless steel (tanks and condencer) ",8]
[10," Water bath\nSupply and install WATER BATH with 27 Liter Capacity Independent  high-limit thermostat provides over-temperature protection. double walled metallic construction ,exterior is stainless steel,18/10 stainless steel tub. Electric heating throat resistance covered  in  stainless steel AISI 316 regulation of temp.by means of digital thermostat PID self adjusting ,2800 w power.        Setting temp.range 10°c to 95°c with additional boilig  stage until 100°c , working temp. range from 5°cabove ambient to nominal temp. (maximum temp ), over heat safety device DIN 12877 thermal classes 1 standard.class2 special  equipment ,integral timer for digital programme time selection, the temp. is measured using a Pt 100 sensor with accuracy of(0.1°c for set temp. and ±0.1°c for temp. fluctuation ), the heater is installed outside to avoid the problems due to dirt lime residue.",8]
[11,"  Autoclaves class ( B )\nTable top horizontal Autoclaves class ( B ) with pre vacuumed stage, vacuumed and heat final dry stage , supplied tray and bag support , bacteriological filter, water distilled tank filled , tray handle extractor, perforated trays in stainless steel ,tray support water tank emptying tap , ventilation bacteriological filter , microprocessor and LED screen, manometer, main switch , chamber capacity  (17 - 23 )L  ",8]
[11.1,"On-site installation "]
[12,"Hot air oven\nMaximal temperature = 200°C\nMicroprocessor PID controller, LCD\ndisplay Timer up to 999h59m, Real-Time Clock,\nProgramming: 50 Programs with 15 steps\n(Temperature, Gradient, Time, Fan Speed), RS 232 Interface, Automatic Restart after Power Cut-Off, one shelf. Internal chamber volume 40Lt.",8]
[12.1,"On-site installation "]
[]
[13," Dental chair  (computerized system) or ( pneumatic)with : dental chair, treatment unit, water unit with (Vacuum system) dry (suction dry line have a container and connect to the   main drain of the chair suction dry line connect to out door, automatic separator , assistant unit and instrument tray) contains: Distilled water container ,Two (solid Titanium body) air turbines with adapters  with fiber optic light ,push  button ,Micromotor with (Titanium angle & straight hand piece) with fiber optic light ,push button , Ultrasound scalar with fiber optic light , Two triple syringe,  Light cure & projector unit LED with sensor switch , Suction  with metal cabinet. unit complete (Durr Dent or Cattani) and saliva ejector Instrument tray with arm., Dentist chair & assistant chair. All air, water, suction, drain piping and fitting , connection , filtering ,main water , special filter, all wiring and connection electrical  , rotor Turbine (10) ,angle head complete with rotor &shaft(10) , extra LED light for projector (2 psc) & LED light cure(2 psc),with all accessories Sirona c8  or  Belmont ClestaII / Adec or Planmeca or equivalent specifications     Oilless compressor(Durr Dent ,jun- air,or Cattani)Air Drier system.with  Compressor metal cabinet                             ",8]
[14,"Amalgamator  Capsule unit ,  protection  cover , Electronically controlled display , permanent memory , compact design  Degussa ,W&H ,or equivalent specifications",8]
[15,"Dental x- ray film processor can be used in manual mode for quick chair -side developeing , full automatic  on and off operation , holds up to 3 film holders at one time, auto \"power- on, power -down'drier fan> ",8]
[16,"Dental x-ray unit \nClassification: Electromedical, Class I, Type B   \nPower supply: 220V±10%,50Hz ,1kVA   \nTube voltage: 60KVp   \nTube current: 8mA   \nFocus size: 1.5mm   \nTotal filtration: 2.5mmAL   \nExposure time: 0.2 to 4 sec   \nLeak radiation: Outside one meter≤0.002mGy/h   \n(National standard: 0.25 mGy/h)   \nTotal weight: 35kg   ",8]
[16.100000000000001,"On-site installation and training  "]
[]
[17," \nConventional X - Ray  Machine\n 500mA ,simple bucky , chest stand, consisting of : X-ray High Voltage Generator, 50kW, Maximum inverter frequency: 50kHz \n- Radiographic kV range : 40 - 150kV at 1kV step, Radiographic mA range: 10 - 500mA, Radiographic mAs range: 0.5 - 500mAs, APR : Max. 245 programs\nFloor-mounted tube support : Floor-to-wall/ceiling tube mount, Longitudinal movement: 2500mm , Transverse movement: 250mm , Vertical X-ray tube stroke, 400-1950mm, X-ray tube rotation on horizontal axis, +-180deg\nCollimator, Multi-leaf collimator\nX-ray tube, 200kHU High-speed X-ray tube, Focal spot: 0.6/1.2mm, Built-in High-speed starter\nHigh tension cable, 14mt\nElevating bucky table, Loading capacity: 200kg (440lbs), Table top longitudinal slide: 1100mm, Table top transverse slide: 250mm, Table top lifting: 535 - 850mm, Table top size: 2200x810mm, Cassette: 6.5\"x8.5\", to 14\"x17\" vertical &horizontal, Grid: 40lines/cm, 12:1 \n  Wall bucky stand \nBucky top to floor distance: 627-2147mm, Chin support to film top distance: 32mm, Bucky vertical dimension: 552mm, Cassette: 6.5\"x8.5\" to 14\"x17\" vertical & horizontal, Grid: 40lines/cm, 12:1 \n(UD150L-40E ) shimadzu , or equivalent specification\n",8]
[17.100000000000001,"On-site installation and training  for X Ray machine"]
[18,"Digital CR Reader \n floor stand  with keyboard holder \ninclude PC, monitor, keyboard,  or barcode reader option \nPlates per hour (70 cassette/hr)\n Cassette sizes: ( 8\" x10\") ,(10\"x12\"),14\")x 14”),(14\"x 17”)\nTime to first image :high-speed scan mode\nGrayscale resolution : Acquisition: 16 bits per pixel\nDisplay: 12 bits per pixel\nMonitors:\n17” fat panel monitor, 1280 x 1024\n19” fat panel touch screen monitor,  \n1280 x 1024\nPower : 200/230V AC  50/60Hz \nUPS ,table and roof. \nWith required soft ware and accessories",8]
[19,"X-Ray viewer\nLED Technology \nSingle Panel\nVaraible intensity: 4,000 10,000 lux\nDimensions: 500 x 530 x 45 mm\nInput voltage : 240 V 50 Hz\nWeight: 6.5 Kg",16]
["Total :"]
[]
[]
[null,"Sub Total Amount for Ramadi Sub PHC:"]
[null,"Sub Total Amount for Ramadi Main PHC:"]
[null,"Total:"]